## Slide 1: Nitrosamine impurities analysis using Thermo Scientific™ Orbitrap Exploris™ 120 MS
Hao Yang, Ph.D.
Sr. Application Scientist
Biopharma and Pharma Vertical Marketing

## Slide 2: Application: Nitrosamine impurities analysis
Liquid chromatography high resolution mass spectrometry (LC-HRMS) method for NAs detection and quantitation
Nitrosamines (NAs) in APIs and drug products
  Classified as genotoxic impurities, proven as probable human carcinogen
  Detected elevated levels of NDMA and other impurities in several drug products, announced several recalls
  US FDA published validated GC/LC-MS/MS and GC/LC-HRAM methods, and interim acceptable limits for several NAs
  US FDA recently released guidance for industry to 1) conduct risk assessments on approved or marketed drugs and pending applications, and 2) implement control strategy to reduce and prevent formation of NAs in all API and drug products
Goals
  1 single LC-HRMS method for detection and quantitation of 9 NAs
  High selectivity and sensitivity
  Robustness and reproducibility
  Compliant software for data collection and processing
Notes: Nitrite are common impurities in excipient (drug formulations) at ppm level

## Slide 3: Application: Nitrosamine impurities analysis
Thermo Scientific™ VanquishTM LC and Acclaim Polar Advantage II column, 2.1 x 100mm, 2.2µm
  Excellent retention time and injection reproducibility
  Good peak shape for target NAs
  No carryover for target NAs, except NDBA
Thermo Scientific™ Orbitrap ExplorisTM 120 MS
  Max 120K resolution at m/z 200
  Sub ppm mass accuracy with EASY-IC
  LLOQ ≤ 0.017 ppm for target NAs, in APCI mode for both neat and excipient standards
Thermo Scientific™ Chromeleon™ CDS 7.2.10
  Single platform for MS and chromatography
  Fully integrated with instrument control, data collection, processing and reporting
  21 CFR Part 11 compliant ready with full instrument and data audit trail

## Slide 4: Application: Nitrosamine impurities analysis
Vanquish LC and Acclaim Polar Advantage II column, 2.1 x 100mm, 2.2µm
  Excellent retention time and injection reproducibility
  Good peak shape across all NAs
  No carryover for target NAs, except NDBA
N = 50, Neat, APCI
Average RT: 0.87 min
RT reproducibility: 0.69%
Average Peak area: 11700
Peak area reproducibility: 2.1%
Robust and reproducible injections

## Slide 5: Application: Nitrosamine impurities analysis
Vanquish LC and Acclaim Polar Advantage II column, 2.1 x 100mm, 2.2µm
  Excellent retention time and injection reproducibility
  Good peak shape across all NAs
  Minimum carryover for target NAs, except NDBA
XIC of NAs at 50ng/ml, Excipient, APCI
Carryover, Excipient, APCI
~0.2%
Good separation with minimal carryover for excipient standards

## Slide 6: Application: Nitrosamine impurities analysis
Orbitrap Exploris 120 MS
  Max 120K resolution at m/z 200
  Sub ppm mass accuracy with Thermo Scientific™ EASY-IC™ ion source
  LLOQ ≤ 0.017 ppm for target NAs, in APCI mode for both neat and excipient standards
High sensitivity, could achieve LLOQ ≤ 0.017 ppm for target NAs in both neat and excipient matrix
 | Matrix | LOD (ng/ml) | LLOQ (ng/ml) | PPM* | Linearity
NDMA | Neat | 0.2 | 0.2 | 0.0068 | 0.2 – 50
 | Excipient | 0.2 | 0.2 | 0.0068 | 0.2 – 50
NMEA | Neat | 0.2 | 0.2 | 0.0068 | 0.2 – 50
 | Excipient | 0.2 | 0.2 | 0.0068 | 0.2 – 50
NPYR | Neat | 0.1 | 0.2 | 0.0068 | 0.2 – 50
 | Excipient | 0.2 | 0.2 | 0.0068 | 0.2 – 50
NDEA | Neat | 0.1 | 0.1 | 0.0034 | 0.1 – 50
 | Excipient | 0.1 | 0.1 | 0.0034 | 0.1 – 50
NPIP | Neat | 0.1 | 0.1 | 0.0034 | 0.1 – 50
 | Excipient | 0.2 | 0.2 | 0.0034 | 0.1 – 50
NEIPA | Neat | 0.5 | 0.5 | 0.017 | 0.5 – 50
 | Excipient | 0.5 | 0.5 | 0.017 | 0.5 – 50
NDIPA | Neat | 0.1 | 0.1 | 0.0034 | 0.1 – 50
 | Excipient | 0.1 | 0.1 | 0.0034 | 0.1 – 50
NDPA | Neat | 0.1 | 0.1 | 0.0034 | 0.1 – 50
 | Excipient | 0.1 | 0.1 | 0.0034 | 0.1 – 50
NDBA | Neat | 0.1 | 0.5 | 0.017 | 0.5 – 50
 | Excipient | 0.1 | 0.5 | 0.017 | 0.5 – 50
* PPM is calculated based on 30mg/ml of drug substance and product extract

## Slide 7: Application: Nitrosamine impurities analysis
Orbitrap Exploris 120 MS
  Max 120K resolution at m/z 200
  Sub ppm mass accuracy with EASY-IC ion source
  LLOQ ≤ 0.017 ppm for target NAs, in APCI mode for both neat and excipient standards
Accuracy and precision for excipient standards at 0.017 ppm
* Accuracy and precision based on 5 replicate injections, data were processed by linear with offset, 1/x weighting
 | % Diff | Precision
NDMA | 0.7 | 7.0
NMEA | -5.2 | 4.2
NPYR | -7.7 | 2.7
NDEA | -5.4 | 3.4
NPIP | -2.2 | 5.1
NEIPA | -6.7 | 7.8
NDIPA | -3.8 | 2.9
NDPA | -4.3 | 3.0
NDBA | -3.3 | 4.1

## Slide 8: Application: Nitrosamine impurities analysis
Orbitrap Exploris 120 MS
  Max 120K resolution at m/z 200
  Sub ppm mass accuracy with EASY-IC ion source
  LLOQ ≤ 0.017 ppm for target NAs, in APCI mode for both neat and excipient standards
High selectivity can be achieved with maximum 120K resolution and sub ppm mass accuracy
NDMA
0.0 ppm
DMF, 15N isotope
0.4 ppm
DMF, 13C isotope
0.0 ppm
N,N-dimethylformamide (DMF)15N isotope and NDMA are baseline resolved at 120K resolution

## Slide 9: Application: Nitrosamine impurities analysis
Orbitrap Exploris 120 MS
  Max 120K resolution at m/z 200
  Sub ppm mass accuracy with EASY-IC ion source
  LLOQ ≤ 0.017 ppm for target NAs, in APCI mode for both neat and excipient standards
High selectivity can be achieved with maximum 120K resolution and sub ppm mass accuracy
With inadequate mass tolerance setting, the presence of DMF in Ranitidine tablet extract could cause overestimation of NDMA
NDMA with 25ppm mass tolerance
Measured: 2.8 ng/ml or 0.093ppm
NDMA with 3ppm mass tolerance
Measured: 2.4 ng/ml or 0.08ppm

## Slide 10: Application: Nitrosamine impurities analysis
Instrument control, data processing, and reporting all in single package
Chromeleon CDS 7.2.10
  Single platform for MS and chromatography
  Fully integrated with instrument control, data collection, processing and reporting
  21 CFR Part 11 compliant ready with full instrument and data audit trail

## Slide 11: A complete Thermo Fisher Scientific solution
A robust and reproducible LC-HRMS method for detection and quantitation of 9 NAs in Ranitidine tablet
High resolution, minimum of 60K, is required to separate DMF isotope interferent peaks to avoid potential isobaric interferences
Adequate mass accuracy and tolerance setting is essential for reliable and accurate quantitation of NDMA in presence of co-eluting DMF in tablet extract
Achieved superior sensitivity with APCI, LLOQ ≤ 0.017 ppm for all NAs tested (FDA proposed limit is 0.03 ppm)
Demonstrated the use of Chromeleon 7.2.10 compliant ready CDS software for instrument control, data acquisition, processing, and reporting

## Slide 12: Backup slides
Recovery and reproducibility of NAs from extraction process
LLOQ of NAs using HESI mode for both neat and excipient standards
LLOQ comparison between HESI vs APCI mode for target NAs

## Slide 13: Application: Nitrosamine impurities analysis
NAs recovery and reproducibility from the extraction process
Crush tablet and extract with methanol to give target concentration of 30mg/ml API
Mechanical shaker for 40 minutes
Centrifuge @ 4500 rpm for 15 minutes
Filter the supernatant with 0.22 µm PVDF filter
 | 2ng/ml |  | 5ng/ml | 
 | % Recovery | % RSD | % Recovery | % RSD
NDMA | 99 | 8.4 | 97 | 7.2
NMEA | 93 | 9.0 | 97 | 4.5
NPYR | 101 | 3.6 | 99 | 5.3
NMBA | 94 | 13.3 | 94 | 8.8
NDEA | 95 | 3.0 | 97 | 3.4
NPIP | 97 | 1.6 | 96 | 3.8
NEIPA | 116 | 10.5 | 98 | 7.0
NDIPA | 95 | 3.1 | 100 | 2.9
NDPA | 105 | 2.2 | 104 | 1.8
NDBA | 101 | 2.8 | 100 | 4.9

## Slide 14: Application: Nitrosamine impurities analysis
Orbitrap Exploris 120 MS
  Max 120K resolution at m/z 200
  Sub ppm mass accuracy with EASY-IC ion source
  LLOQ ≤ 0.068 ppm for target NAs in HESI mode for both neat and excipient standards
High sensitivity, LLOQ ≤ 0.068 ppm for target NAs for both neat and excipient standards, HESI mode
 | Matrix | LOD (ng/ml) | LLOQ (ng/ml) | PPM* | Linearity
NDMA | Neat | 0.2 | 0.5 | 0.017 | 0.5 – 50
 | Excipient | 0.5 | 0.5 | 0.017 | 0.5 – 50
NMEA | Neat | 0.5 | 0.5 | 0.017 | 0.5 – 50
 | Excipient | 0.5 | 0.5 | 0.017 | 0.5 – 50
NPYR | Neat | 0.1 | 0.1 | 0.0034 | 0.1 – 50
 | Excipient | 0.1 | 0.2 | 0.0068 | 0.2 – 50
NDEA | Neat | 0.2 | 0.5 | 0.017 | 0.5 – 50
 | Excipient | 0.5 | 0.5 | 0.017 | 0.5 – 50
NPIP | Neat | 0.2 | 0.2 | 0.0068 | 0.2 – 50
 | Excipient | 0.2 | 0.2 | 0.0068 | 0.2 – 50
NEIPA | Neat | 1 | 2 | 0.068 | 2 – 50
 | Excipient | 2 | 2 | 0.068 | 2 – 50
NDIPA | Neat | 0.2 | 0.5 | 0.017 | 0.5 – 50
 | Excipient | 0.2 | 0.5 | 0.017 | 0.5 – 50
NDPA | Neat | 0.2 | 0.2 | 0.0068 | 0.2 – 50
 | Excipient | 0.2 | 0.2 | 0.0068 | 0.2 – 50
NDBA | Neat | 0.1 | 0.5 | 0.017 | 0.5 – 50
 | Excipient | 0.1 | 0.5 | 0.017 | 0.5 – 50
* PPM is calculated based on 30mg/ml of drug substance and product extract
We can achieve less than 0.017 ppm for all NAs except NEIPA in HESI mode

## Slide 15: Application: Nitrosamine impurities analysis
Orbitrap Exploris 120 MS
  Max 120K resolution at m/z 200
  Sub ppm mass accuracy with EASY-IC ion source
  Superior sensitivity can be achieved using both HESI and APCI mode
LLOQ comparison between HESI and APCI for all NAs (Neat only)
 | Mode | LOD (ng/ml) | LLOQ (ng/ml) | PPM* | Factor
NDMA | HESI | 0.2 | 0.5 | 0.017 | -
 | APCI | 0.2 | 0.2 | 0.0068 | 2.5x
NMEA | HESI | 0.5 | 0.5 | 0.017 | -
 | APCI | 0.2 | 0.2 | 0.0068 | 2.5x
NPYR | HESI | 0.1 | 0.1 | 0.0034 | -
 | APCI | 0.1 | 0.2 | 0.0068 | 1/2x
NDEA | HESI | 0.2 | 0.5 | 0.017 | -
 | APCI | 0.1 | 0.1 | 0.0034 | 5x
NPIP | HESI | 0.2 | 0.2 | 0.0068 | -
 | APCI | 0.1 | 0.1 | 0.0034 | 2x
NEIPA | HESI | 1 | 2 | 0.068 | -
 | APCI | 0.5 | 0.5 | 0.017 | 4x
NDIPA | HESI | 0.2 | 0.5 | 0.017 | -
 | APCI | 0.1 | 0.1 | 0.0034 | 5x
NDPA | HESI | 0.2 | 0.2 | 0.0068 | -
 | APCI | 0.1 | 0.1 | 0.0034 | 2x
NDBA | HESI | 0.1 | 0.5 | 0.017 | -
 | APCI | 0.1 | 0.5 | 0.017 | 1x
* PPM is calculated based on 30mg/ml of drug substance and product extract